INDICATION FOR USE

The ScopeFlow Pure is intended to be used in conjunction with a flexible endoscope to insufflate air/CO2 and supply sterile water during endoscopic procedures. The ScopeFlow Pure is compatible with the different sterile water bottles available. It is intended to be used for 24 hours and then disposed of.

CAUTION

Federal (United States) law restricts this device to sale by or on the order of a physician.

SYMBOL LEGEND :

| Manufacturer | Do not use if pack is opened or damaged | Consult Instructions for Use | Sterilized using Ethylene Oxide | Date of manufacture | Use by date | Product reference number | Lot number |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Keep away from sunlight | Keep dry | Biohazard | Do not re-sterilize | Not made with Natural rubber latex | Not made with DEHP |  |  |

PRODUCT CODES AND DESCRIPTIONS

| Product code | Description | Endoscope compatibility (Gastroscope – Coloscope) |
| --- | --- | --- |
| SFP145 | ScopeFlow Pure™ for Olympus 140 & 240 series endoscopes and higher | Olympus® series 140/ 240, 160/260, 180/280, 190/290 |
| SFP145CO2 | ScopeFlow Pure™ for Olympus 140 & 240 series endoscopes and higher with CO2 | Olympus® series 140/ 240, 160/260, 180/280, 190/290 |
| SFP145CO2S | ScopeFlow Pure™ for Olympus 140 & 240 series endoscopes and higher with CO2 (short) | Olympus® series 140/ 240, 160/260, 180/280, 190/290 |
| SFP1451 HYBRID | ScopeFlow Pure™ for Olympus 140 & 240 series endoscopes with CO2 and EndoStream™ PFE130 | Olympus® series 140/ 240, 160/260, 180/280, 190/290 |
| SFP1452 HYBRID | ScopeFlow Pure™ for Olympus 140 & 240 series endoscopes with CO2 and EndoStream™ PFE230 | Olympus® series 140/ 240, 160/260, 180/280, 190/290 |

DESCRIPTION OF USE

The ScopeFlow Pure functions by using a two parallel tubes configuration to supply air to the water bottle and provide water to the flexible endoscope.

The HYBRID configuration consists of the ScopeFlow Pure and the EndoStream units being used together with one single sterile water bottle.

DIRECTIONS FOR USE

Read all of the instructions for use for the disposable water bottle tubing carefully before use.

1) Open the sterile water bottle

2) Remove ScopeFlow Pure tubing from the sterile pouch and ensure the pinch clamp situated on the tubing is open.

3) Insert the tubing into the sterile water bottle and screw the plastic cap of the ScopeFlow Pure to the sterile water bottle.

4) Peel off the 24h00 sticker from the packaging label. Write the date and time on the sticker before attaching it to the tubing.

5) Connect the distal extremity of the tubing ScopeFlow Pure (endoscope adaptor) to the endoscope air/water connector.

6) Turn on the light source of the processor.

7) If using CO2 insufflation, connect the CO2 insufflator tubing to the connector of the ScopeFlow  Pure insufflation and irrigation tubing/or directly to the CO2 insufflator.

8) If using the Hybrid product, please refer to EndoStream Instructions for Use for connection of irrigation tubing to the endoscope and pump at point 6 – (enclosed)

9) Test prior to endoscopic procedure for water flow, if the flow is slow, check the fitting of the screw cap on the sterile water bottle and tighten if necessary.

10) If the sterile water bottle should be changed during the procedure, consult trust guidelines on the correct technique in handling of the ScopeFlow Pure during the replacement of the sterile water bottle, setup of the device and between patients.

11) When the procedure is completed, close the pinch clamp and remove the ScopeFlow Pure from the endoscope. This will prevent water spillage.

12) Turn off the processor

13) Dispose of the ScopeFlow Pure tubing after 24 hours.

WARNINGS AND PRECAUTIONS

- Always prime the air/water channel prior to inserting the digestive endoscope into the patient.

- The manufacturer did not design this device to be reprocessed or used for more than 24 hours once the packaging is opened, and therefore cannot verify that reprocessing can clean and/or sterilize or maintain the structural integrity of the device to ensure patient and/or user safety.  A traceability sticker is provided.

- The ScopeFlow Pure tubing is intended to be used for 24 hours after the packaging is opened. The product should be discarded within this time period regardless of numbers of use.

- Do not attach or reattach the ScopeFlow Pure to an endoscope which has not been reprocessed.

CONTRAINDICATIONS

The contraindications are those which are specific to all endoscopic procedures.

The disposable water bottle tubings should not be re-sterilized.  Only to be used with Sterile water bottles.

STORAGE / PACKAGING

Storage : In order to prevent any damage, products should be stored in their original packaging, away from sunlight in a cool dry place.

Packaging : The disposable water bottle tubings, ScopeFlow Pure, are supplied in an individual sterile package in a box of 10 units.

PRODUCT DISPOSAL

After use, this product may be a potential biohazard. Handle and dispose of in accordance with accepted medical practice.